MODEL NO：UCA-N

BSC-N

UCA-N                     BSC-N

Recharge

1.  when to recharge

- After turning on the scale or in usage, the figures keep shinning, which means the power of battery is not full. It is time to recharge. If it need to use, it may continue about ten hours.

When the message “LoBAt” appears in the window of piece weight, which indicate the power of battery is exhausted. Please turn off the power switch immediately and begin to recharge the battery to prolong its life.

2.  Recharge Notes

It will take 6-8 hours to full the battery.

Maintenance

- Please take good care of the scale for it’s a precision

instrument.

- Place an object on the pan softly in less to destroy load cell

for the power of moment hit is too heavy.

- In order to keep the scale in formal; please recharge the

battery in time.

-8-

You have purchased a quality precision weighing instrument that requires handling with care. Read entire contents of this Operation Manual prior to operating your new instrument.

# Disclaimer Notice

Calibrate your instrument using reference weights of the appropriate tolerance (class).

An instrument can be no more accurate than the standard to which it has been compared.

For assistance in the selection of reference weights, please contact the factory.

Caution: Changes or modifications not expressly approved by the manufacturer could void the user’s authority to operate this equipment.

## Introduction

Thank you for choosing one of our instruments. Your instrument is designed and manufactured to the most rigorous standards in order to give you years of service. First, check the contents of the shipping carton. You should find the following:

*Manual    *Instrument   *Power cable

Next, follow the instructions for installing your instrument.

Now you are ready to begin using your instrument. To take advantage of its many features, carefully read your operating manual. It contains step-by-step procedures, examples, and other vital information.

Warning: Use of this product in a manner not specified by the manufacturer may impair any safety protection provided by the equipment.

-1-

### Specification

1. UCA-N

| Capacity | 3Kg | 6Kg | 15Kg | 30Kg |
| --- | --- | --- | --- | --- |
| Division-A | 0.5g | 1g | 2g | 5g |
| Division-B | 0.2g | 0.5g | 1g | 2g |
| Division-C | 0.1g | 0.2g | 0.5g | 1g |
| Mini. Cap. | 10g/4g/2g | 20g/10g/4g | 40g/20g/10g | 100g/40g/20g |
| Pan size | 285 x 240 mm |
| Power | AC: 110V/220V; DC:4V 4Ah |

2. BSC-N

| Capacity | 60Kg | 150Kg | 300Kg |
| --- | --- | --- | --- |
| Division-A | 10g | 20g | 100g |
| Division-B | 5g | 10g | 50g |
| Division-C | 2g | 5g | ---- |
| Mini. Cap. | 200g/100g/400g | 400g/200g/100g | 2kg/1kg |
| Pan size | 400 x 500 mm (size can be changed as customers’ request) |
| Power | AC: 110V/220V; DC:4V 4Ah |

### Preparation

This product is intended for indoor use.

- Select a suitable work area.

- Work area should be relatively free from drafts and vibrations.

- Work surface should be level and rigid.

- Do not locate near magnetic materials or equipment/instruments which use magnets in their design.

- Avoid areas which have variations in room temperatures or have excessive room temperatures. Room temperatures above 40°C or 0°C could affect instrument operation and accuracy.

-2-

| Display | Option | Parameter |
| --- | --- | --- |
| diu-F01 | Division | 1 2 5 10 20 50 |
| Pnt-F02 | Decimal point | 0.0 0.00 0.000 0.0000 |
| CAP-F03 | Capacity | If cap is 30kg, 30.000 |
| LNE-CAL-F04 | Calibration | only use for distributor |
| tra-F05 | Zero-tracking | 0.5 1.0 1.5 2.0 2.5 3.0 3.5 4.0 |
| Usr-F06 | Quick calibration | See below |
| Lit-F07 | Backlight switch | Au: light during operation, press the key ‘3’ to set backlight time. On: always light |
| Def-F08 | Unit defult | U1 (kg) U2 (g) U3(Lb) |
| Bud-F09 | Baud rate | 2400 4800 9600* 19200 （9600,for connect to printer and PC） |
| Ode-F10 | Print | 0 (no) 1 (manual) 2 (automatical) 3(continuous) 4 (ULP manual) 5 (ULP auto) |
| Upt-F11 | Piece weight change | On(if you want to get more accuracy weight, you can choose this ) off |

Calibration

Usr-F06 (Quick calibration)

When it display

Press the key

It display

Put any STANDARD WEIGHT on the pan, say 2kg weight. After a little while,

it will display

the scale will finish calibration.

-7-

- Alarm functions

- Alarm setting

This function is designed for packing purpose. For example, if the operator wish to count 100 pieces for every package, he can set the upper limit as 100 pieces.

a. Press the        key.

b. Press the key        ,        ,        one by one.

c. Press the       key, to return normal counting function.

After such setting, if total quantity reached 100 or over, scale will

sound bi-bi-bi---, and shows –0.t.y--

- Alarm release

a. Press the key  .     (

- Connect with ULP printer, choose print form.

Turn on the scale, Hold the key       ,till it display [No. XX], press numberic keys to choose print form, and then press the key

to confirm.

Programme

Turn on the scale, and input password “7 5 9”. The display will shows

Press       key for entering setting and confirm setting.

Press          key for setting and goes to next step. The

programme sequence as follows:

-6-

##### Function keys

On/ Off Switch         Turns instrument On or Off.

Numeric keys. Used for setting numeric data for tare weight. sample number, sample weight, or limit number of checking.

- Captures a new center of zero.

- Cancel the memorized data in total mode.

Reduce gross weight on pan as tare weight.

Used for set the decimal position of the tare weight, sample weight.

Used for canceling the numeric setting data.

Used for parts counting.

Used for the alternation of changing normal

counting and check operation.

Used for accumulating the sub-total data.

Unit conversion key.

Used for check quantity

Function key.

-3-

### Operation

- Getting Started

Turn the instrument on by pressing the power switch on.

The display will down count from 999999… while the unit is being updated.

Allow the unit to warm-up for 30 minutes.

- Taring (zeroing)

All models have taring (zeroing) capabilities up to their total weight capacity.

If sample come with receptacle, but you just need to know weight of sample. Please use TARE key.

For reducing the tare weight from the scale, there are two methods, one is press button tare and the other is keyboard tare.

- Press button tare

Place the sample receptacle on pan and then press the key       and then will point to and the WEIGHT display

shows 0.

- Keyboard tare

- Using the numeric and decimal keys to set the known tare weight data into the scale. The data is displayed in the PIECE WEIGHT display with flickering.

- Press       key.

- Set tare weight displayed in the WEIGHT display is cleared to zero. Place sample with receptacle, the WEIGHT display will show the weight of the sample.

- Clearing the previous tare weight. Remove receptacle from the pan then press       key, and the WEIGHT display returns to zero.

-4-

- Counting Function

- PIECE COUNT setting(e.g.)

  - Count the desired sample pieces(e.g.:20 pieces)and put on the pan.(sample size can be changeable)

b. Input the quantity of samples and Press the key

c. The averaged piece weight will be

calculated automatically.

Now you can count bulks.

- PIECE WEIGHT setting

(In case of piece weight is already known.)

- Set the piece weight data by the numeric keys and decimal point key. The number to be shown in PIECE WEIGHT display with flickering.

- Press the       key, the piece weight setting to be done.

- Press the        key to cancel the piece weight setting.

- PIECE WEIGHT alarm

will point to LS and AS; when the averaged piece weight or set piece weight is not enough for accurate counting operation. Operator still can use the scale, but counting error may occur.

D.  Accumulation function

1. Press      key can be accumulating the sub-total piece counts one time. The next accumulation must be return to zero and got new piece counts.

2.Press       key can be used for alternation of changing normal counting and memory data recalling.

3.The       key can be used to delete all data in Grand-Total.

-5-

# Text boxes

9

0

ZERO

TARE

·

CE

SAMPLE

SAMPLE

M+

U

P

M+

CHECK

diU

M+

CE

TARE

TARE

TARE

WT SET

1

WTSET

0

5.0

0.0

20 20100100

5.0 100100

100.0

0

20 100100

100.0

SAMPLE

100.0

0

F01

0

CHECK

CHECK

1

0

0

CHECK

·

0.000

-CAL

USr

F06

CHECK

CHECK

·

MC

0

PUt

AU-CAL

0

0

2.000

Note:

- Please recharge the scale before first operation.

- Please warm-up the scale over 5 minutes.